WinX/32 User Programming –Description for Visual Basic

Princeton Instruments

19 May 1999

Table Of Contents

Overview	1

Beginning in Visual Basic	2

Things to Note:	4

Calling Syntax	4

Getting Data	5

Error Handling	5

List of Functions	5

Winx32App functions	6

DocFile functions	7

DocFiles functions	9

Doc Window functions	10

Doc Windows functions	11

Print Window functions	12

Print Windows functions	13

Experiment Setup functions	14

ValidRange object	17

CalibObj object	17

DocInfo object	17

CursorObj object	18

Parameter lists	19

DM_CMD (DocFile Parameters)	19

DI_CMD (DocWindow Parameters)	21

EXP_CMD (Expsetup Parameters)	23

Creating a Snap-In DLL	27

How to Convert a Visual Basic “Exe” project to be a “Snap-In”	28

Comparison of Visual C++ and Visual Basic Interfaces to WinX/32	29

Index	31

WinX/32 User Programming (Automation)

Overview

WinX/32 Automation provides a way to write programs to control the Winview and Winspec family of PI applications (hereafter refered to as WinX/32). This ability was provided to users of the previous generation of PI software via Macro Basic. Automation allows programs to be written in Visual Basic, Visual C++, or any other application language which supports the Microsoft Windows Automation standard, such as the macro languages in Microsoft Excel and Microsoft Word. (Princeton Instruments technical support will only support Microsoft Visual Basic, Microsoft Excel, and Microsoft Word.)

Each application that supports Automation provides entry points, called interfaces, which can be called by an outside program. Windows uses a file called a type library to provide information about what interfaces an application provides. The Windows registry stores the location of the type library. When WinX/32 with Automation support is installed, the Winx32 type library is registered. This means that “Princeton Instruments WinX/32 Type Library 2.0” will show up in Windows’ list of automation support providers. Visual Basic, Visual C++, and other programs provide a way to look at the list and choose which applications the user’s program will call. In Visual Basic the facility is called the References browser.

Automation programs may be written either as stand-alone programs (executables) or as dynamic link libraries (DLLs.) One advantage of a DLL user program is that WinX/32 will automatically show a menu item and a toolbar button for it. DLL user programs that are integrated with WinX/32 this way are called “Snap-Ins”. The easiest way to produce a program is to write a Visual Basic executable. Visual Basic Snap-Ins require additional effort to produce, and have some limitations: they always run as modal, so there is no possibility of interacting with the WinX program while a VB Snap-In is running; and Visual Basic Snap-Ins cannot be debugged using the Visual Basic debugger.

Microsoft supplies three versions of Visual Basic: the Learning Edition, the Professional Edition, and the Enterprise Edition. You must have the Professional or Enterprise edition in order to produce a DLL. The Learning Edition can produce a compiled executable program, but it may not run as fast as one compiled with the Professional or Enterprise editions. The examples below were produced using the Professional Edition.

Once your program is written, you can run it by clicking on the Run button in Visual Basic. This is a good way to debug your program. When you are fairly confident that the program is bug-free, you can compile it as an executable. The program can then be run in the normal way from an explorer window, or from the Execute Macro menu item in WinX/32. The Execute Macro menu item remembers the last 8 macros, so they can be selected quickly from a list.

A Visual Basic executable project can be converted to run as a Snap-In DLL. You must know how to add a resource file to the project; also, Visual Basic Snap-Ins have limitations which stand-alone programs do not have. Princeton Instruments also supplies a Visual Basic Snap-In Wizard, which will set up a new Snap-In DLL project.

The macro languages of Microsoft programs such as Excel or Word also use the Visual Basic syntax, and can be used to write programs which call WinX32 functions. These macro programs run in interpreted mode, so things like For...Next loops will be slower than with a compiled program. However, if your macro mainly consists of calls to built-in functions of these applications, the overall speed of your macro may still be very good.

# Beginning in Visual Basic

The first example here will describe how to create a Visual Basic executable program which can call functions in WinX32. Start a new project of the type "Standard Executable". Visual Basic will create a project with a standard form or dialog. Your VB program will display this form when it first starts to run. Usually you will write your VB program by drawing controls on this form, and by specifying what happens when a user interacts with these controls. If you want to make the Basic program begin some action immediately when it starts, without waiting for user interaction, put the code in the Load function for the start-up form, or else set a function or subroutine as the start-up object instead of the form.

To gain access to the WinX32 functions, go to the References browser (in the Project menu) and check the item “Princeton Instruments WinX/32 2.0 Type Library” (see the figure on the previous page). If this item is not present, then WinX/32 was not installed properly. You may also want to check other items that are part of the Princeton Instruments collection of type libraries, if they are available. For instance, the type library labeled “Princeton Instruments’ PIXBM32” provides additional facilities for data acquisition and processing of images or spectra.

Once the correct references have been selected, the Visual Basic Object Browser (under the View menu) will show all of the functions and parameters available in the selected type libraries.  To see the functions in the WinX32 type library, use the drop-down box at the top of the Object Browser to select “Winx32Lib”; it defaults to “All”. For each function or parameter, a short help string or "hint" is available and will be displayed at the bottom of the Object Browser window when you select it. Part of this hint will also appear when you are typing the function or parameter into your Visual Basic program. You can leave this window open while you are writing your VB program to show what interfaces and functions are available to you.

All of the WinX32 functions must be called via an object from the type library. An object can be created in Visual Basic using the following syntax:

Dim <object name> as New Winx32Lib.<object type>

For example:

Dim WApp as New Winx32Lib.Winx32App

You can omit the "Winx32Lib" part of these statements if you are sure no other type library also supplies an object of that type. For instance, if some other program also had an interface called "DocFile", and you selected both that program and WinX32 in the References browser, then VB  would have no way of knowing which DocFile object to create.

Once the object is created, you can call all of the functions of the associated interface as member functions (methods):

WApp.ShowDemoBox("This is a demo box”)

Create a button on your new form by clicking the button icon on the VB toolbox and then drawing a box on the form. When you are done you should see something like this:

If you have trouble figuring out how to place a button on a form, or want more information, look in the Visual Basic Help Contents for "Designing a Form".

To get to the code which will execute when a user clicks on the new button, double-click on it with the mouse. You will be taken to a code window looking like the one below. Into that window, after the line which says

Private Sub Command1_Click()

type the two lines already given, that is:

Dim WApp as New Winx32Lib.Winx32App

WApp.ShowDemoBox("This is a demo box”)

The result should look like the figure below:

Now just click on the Start button in the VB menu bar, or else choose the Start item under Run. You should see your new form come up. Click on the button. If a WinX32 program wasn't running yet, it will start up at that point, and display a message box with the contents "Hello World!"

Next try giving your program the ability to collect an image or spectrum. Stop the running VB program, and replace the code in Command1_Click() with the following:

Private Sub Command1_Click()

Dim theFrame As Variant

Dim ex As New ExpSetup

Dim dx As DocFile

Dim bRes As Integer

If ex.Start(dx) Then ‘ make sure detector started

While ex.GetParam(EXP_RUNNING, bRes)

DoEvents ‘ wait for detector to finish

Wend

dx.GetFrame 1, theFrame ‘ retrieve 1 image

For i = LBound(theFrame, 1) To 5 ‘ for a few points

s = s & theFrame(i, 0) & ", "

Next i

MsgBox s ‘ display a few points

End If

End Sub

This little routine will start an Acquisition, wait until it is done, retrieve the first frame of data, and display a few values from it.

Finally try doing something with the data and putting it back into WinX. Replace the code in Command1_Click with the following example that takes an image, subtracts an offset, then puts the image back into WinX:

Private Sub Command1_Click()

‘ This example collects a frame of data as defined in

‘ experiment setup. It then subtracts a hardcoded dc

‘ offset from the image and then puts the image back

‘ into WinView and updates the display

Dim theFrame As Variant  ‘ Array that holds the image

Dim ex As New ExpSetup ‘ Experiment object

Dim dx As DocFile      ‘ Document/file object

Dim bRes As Integer

If (ex.Start(dx)) Then ‘ Start the experiment, check success

‘ Wait for experiment to complete. The EXP_RUNNING

‘ param returns 1 if experiment is still running

‘ and 0 if the experiment is finished.

While ex.GetParam(EXP_RUNNING, bRes)

DoEvents ‘ Let other events get processed.

Wend

dx.GetFrame 1, theFrame ‘ retrieve 1 image

Xdim = dx.GetParam(DM_XDIM) ‘ get the document x dimension

‘ For dimensions you can either use the DM_XDIM or you can

‘ get the following bounds for the array.

‘ l1 = Lbound(theFrame, 1)

‘ u1 = Ubound(theFrame, 1)

Ydim = dx.GetParam(DM_YDIM) ‘ get the document y dimension

‘ For dimensions you can either use the DM_YDIM or you can

‘ get the following bounds for the array.

‘ l1 = Lbound(theFrame, 2)

‘ u1 = Ubound(theFrame, 2)

dcOffset = 10 ‘ subtract a value from the image

For i = 0 to (Ydim – 1)   ‘Loop through y dim

For j = 0 to (Xdim – 1) ‘ Loop through x dim

theFrame(j, i) = theFrame(j, i) - dcOffset

Next j

Next I

‘ Put image back to WinX

dx.PutFrame 1, theframe

‘ Update display in WinX for image

‘ Note if this is not done the data will be changed but

‘ the image display will show the old data.

dx.Update

End If

End Sub

At this point, you know enough to automate many common tasks, using stand-alone Visual Basic programs to control WinView or WinSpec. To find out how to create a Snap-In DLL, see the section “Creating a Snap-In DLL”.

## Things to Note:

#### Calling Syntax

Visual Basic differentiates between Functions, which return some value, and Subroutines, which do not. Some of the methods of WinX/32 COM objects are defined as functions. When calling either, you can put the arguments in parentheses, or else just list them after the function name. For example, the function SetParam returns a value to indicate whether it succeeded in setting the designated parameter. If you don’t want to bother checking whether SetParam succeeds (although you really should!) you can call it without assigning the return value to anything:

Dim dx as new DocFile

dx.SetParam(DM_XDIM, vX) ' with parentheses (gives syntax error)

dx.SetParam DM_XDIM, VX  ' without parentheses

The line with parentheses above will generate a syntax error if you enter it as shown. This is because SetParam is defined as a Function which returns a boolean value. VB won't allow you to write the call to a function in the parenthesized form unless you also assign the return value to something; for instance:

bRes = dx.SetParam(DM_XDIM, vX) ' with parentheses; this will work

In this case, VB creates a variable of type Boolean and assigns the return value of SetParam to it.

The function GetParam returns the value of the designated parameter as its return value, e.g.:

Xdim = dx.GetParam(DM_XDIM) ' get the document x dimension

so you must use the syntax with parentheses.

Subroutines with multiple parameters are also sensitive to the presence or absence of parentheses. In this case, if you use parentheses, you have to proceed the Subroutine call with the Call keyword:

dw.SetPosition top, left, right, bottom ' without parentheses

Call dw.SetPosition(top, left, right, bottom) ' with parentheses

More difficulties arise with Functions which take a variable number of arguments; for instance, the Print function of PrintWindow. The best way to call this function is to use the form with parentheses, and assign the return value:

bRes = px.Print(1, 1, 0, "This is ", 22, " characters") ' this will work

#### Getting Data

When you use the ExpSetup object's Start or StartFocus function to begin collecting data, WinX/32 launches a separate thread to run the acquisition task. This means that your VB program can continue to run while data is being collected. Therefore, you may need to cause the program to wait until data is actually available before proceeding. Visual Basic provides the DoEvents function to allow a program to wait, while allowing the user to interact with the program in such ways as moving the window, clicking on a Stop button, etc. The code fragment below illustrates this process:

Dim dx as DocFile ' a document for the data

Dim vx as Variant ' a local array for the data

Dim ex as New ExpSetup ' to start data collection

Ex.Start(dx) ' take one set of frames and stop

While ex.GetParam(EXP_RUNNING) ' wait for data to be acquired

DoEvents ' allow Windows to process messages

Wend

dx.GetFrame 1, vx ' now get data

#### Error Handling

Be sure to take note of the error-handling code on page 25.

List of Functions

What follows is a list of the functions that WinX32 supports via the Automation interface. Each section contains the functions for one interface, which corresponds to a Visual Basic object.

In addition to these functions, other functions are available in the EasyDLL and PI Basic Interface libraries. These libraries provide additional capabilities for controlling the detectors, collecting data, and processing that data.

# Winx32App functions

To use these functions, the user program must create an object of type Winx32App and call the functions through the object.

- Sub ShowDemoBox(Text as String)

- Displays the text in pText via a simple message box.

- Example:

- Dim ax as new Winx32App

- ax.ShowDemoBox "This is my text"

- Function GetAppWnd() as Long

- Returns the Win32 window handle of the WinX/32 application's main window. Can be used with SendMessage, ShowWindow, and other Windows API functions.

- Example:

- Public Declare Function ShowWindow Lib "user32" (ByVal hwnd As Long, ByVal nCmdShow As Long) As Long

- Public Const SW_SHOWMAXIMIZED = 3

- Dim ax as new Winx32App

- Dim lx as long

- lx = ax.GetAppWnd

- ' Maximize the WinX application window

- r = ShowWindow(lx, SW_SHOWMAXIMIZED)

- Function GetAppRect(top as long, left as long, bottom as long, right as long) as Boolean

- Retrieves the coordinates of the WinX/32 application's main window. Returns non-zero on error.

- Function CountOpenDocs() as Long

- Returns the number of documents open.

- Function CloseOpenDocs(cDocs as short) as Boolean

- Closes all open document windows. If cDocs is not 0, it will have the number of windows closed on return. Returns 0 if no error.

- Function InfoBox(Boolean bOn) as Boolean

- Opens or closes the info box. Returns TRUE if the box was open before the call.

- Example:

- Dim ax as winx32app

- bRes = ax.InfoBox(TRUE) ' turn on the info box

- Function StatusBar(Boolean bOn) as Boolean

Opens or closes the Status Bar. Returns non-zero if the bar was open before the call.

- Function Version as String

This function returns a string which identifies the version of Winview or Winspec which is running. For instance, version 2.4 of Winspec returns a string like “2.4.0.0”

- Function Flavor as String

This function returns a string which identifies which WinX32 program is running. For instance, Winview/32 returns the string “WinView/32”, and WinSpec/32 returns “WinSpec/32”.

- Sub Hide(Boolean bhide)

This function can hide or show the WinX/32 application and all child windows. If the user did not start WinView or WinSpec before running your program, the first time your program creates a COM object (such as a WinX32App or a DocFile) then WinView or WinSpec will start automatically. When the program detects that it is starting as a result of a COM object call, it will wait an additional 3 seconds before displaying its splash screen or main window. This is to give your program enough time to call the Hide function, which will prevent the window from ever being displayed. You can also call the Hide function after the WinX/32 application is visible. If the Boolean parameter is true, the application will not be visible; if False, if will be visible.

- Function StatusBarMsg(Index as integer, Color as integer, Text as String) as Boolean

Use this function to display a message in the status bar at the bottom of the WinX/32 application. The status bar is divided into 4 panes. The leftmost pane corresponds to an Index of zero. Colors are defined as:

| Note that the WinX/32 application uses the status bar panes for its own purposes, and that the text you place onto the pane may be subsequently overwritten by the application. Black | 1 | Yellow | 7 |
| --- | --- | --- | --- |
| Red | 2 | Dark Gray | 8 |
| Green | 3 | Gray | 9 |
| Blue | 4 | Light Gray | 10 |
| Cyan | 5 | White | 11 |
| Pink | 6 |  |  |

- Sub ShowProgress(Percent as integer)

Use this function to control the progress indicator, which appears on the second pane of the status bar. Normally, this indicator is used to indicate the progress of some process the application is running. The value of Percent should be between zero and 100. A value of 0 makes the progress indicator disappear. Any other value will cause it to appear and display a bar to indicate the value of Percent. Since the application also uses the progress indicator, your setting may be overwritten if data collection or some other process is running.

- Function Quit() As Boolean

Use this function to attempt to close WinX/32. This function acts as if the user clicked on the system close button ([x] in the top right corner of the window). In certain situations (such as the program in data acquisition mode), the user is prompted as to whether or not close the program. If the user declines to close, WinX/32 will remain open. The return value is True if WinX/32 was closed successfully, otherwise the return value is False.

- Function ShowFileOpen(InName As String) As String

Function ShowFileSaveAs(InName As String) As String

Use these functions to show WinX/32’s File Open/Save-As Dialog. The dialog is initialized with InName. If Open/Save was pressed, the function returns the name of the file selected (including the path).  If Cancel was pressed, the function returns an empty String.

# DocFile functions

To use these functions, the user program must create an object of type DocFile and call the functions through the object. A DocFile object has these properties:

Properties:

| Name | Native type | Description |
| --- | --- | --- |
| Title | String | The title in its associated DocWindow (excluding the dimensions) |

- Function Open(Name as String [, xLen as integer, yLen as integer, zLen as integer, datatype as Long, newName as String]) as Boolean

Gets a pointer to an open document. If a name is supplied as the first parameter, then a pointer to an open document is returned, or a document is opened from disk (the document must already exist). The other parameters are optional, if Name is not a null string. If Name is "", Open creates a new document using the values for xLen, yLen, zLen and datatype to determine the size of the document, and newName to determine the name. If newName is "", the new document gets the default name. If xLen, yLen, or zLen are 0, the current controller setup determines the size of the document.

The data type can be as follows:

| 0 | Use detector default | 5 | 8 bit unsigned |
| --- | --- | --- | --- |
| 1 | 16 bit integer | 6 | 16 bit unsigned integer |
| 2 | 32 bit integer | 7 | 32 bit unsigned long |
| 3 | 32 bit single precision floating point |  |  |
| 4 | 64 bit double precision floating point |  |  |

Example:

Dim dx as new WINX32Lib.DocFile

dx.Open("d:\data\Circuit.spe")

- Function OpenNew(Name as String, pInfo as DocInfo) as Boolean

- This is an alternative way to open a data file with the DocFile object. This method uses the DocInfo object to specify information about the new document. In addition to the parameters used in the Open function, such as Xlen, Ylen, and Datatype, the DocInfo object allows you to specify whether to open the file in Append mode and whether to display the file in a window.

- When a data file is opened in Append mode, any old file with the same name is not erased. Instead, the number of frames specified in the pInfo structure is added onto the end of the old file. The parameter DM_LASTFRAMERDY specifies which is the last frame with valid data. When a new file is first opened, it is set to zero. If an old file is appended to, then DM_LASTFRAMERDY is set to the number of frames in the file before the append occurred. The PutFrame and PutStrip functions update the DM_LASTFRAMERDY parameter; However, if you put data into the file using PutPixel, you must update DM_LASTFRAMERDY  yourself.

- For more information, including how to use the bShowWindow member of DocInfo, see the description of DocInfo.

- Function GetWindow() as Object

- Returns a pointer to an object of type DocWindow, which you may use to call the Doc Window interface (display) functions for a window displaying this document.

- Sub GetFrame(frame as integer, buffer as Variant)

- Sub PutFrame(frame as integer, buffer as Variant)

- Sub GetStrip(frame as integer, strip as integer, buffer as Variant)

- Sub PutStrip(frame as integer, strip as integer, buffer as Variant)

- Function GetPixel(frame as integer, strip as integer, pixel as integer) as Variant

- Sub PutPixel(frame as integer, strip as integer, pixel as integer, pixel as Variant)

- GetFrame copies the data from a document into a Visual Basic array. If buffer is an empty Variant, GetFrame creates an array of the proper size and data type and sets buffer to point to it before copying the data.

- PutFrame copies the data from a Visual Basic array into a document. The array should be one created with GetFrame to ensure that the size and data type match.

- GetStrip and PutStrip are similar to GetFrame and PutFrame; they copy one strip to or from the document.

- GetPixel returns a single pixel value from the document. PutPixel sets the value of a single pixel in the document. The pixel value is automatically converted from the type of the Visual Basic variable used to the correct data type for the document.

- Example:

- Dim dx as WINX32Lib.DocFile

- Dim dBuf as Variant

- dx.Open("d:\data\circuit.spe")

- dx.GetFrame 1, dBuf

- ' find the size of the 2-D array returned

- l1 = LBound(dBuf, 1): l2 = LBound(dBuf, 2)

- u1 = UBound(dbuf, 1): u2 = UBound(dBuf, 2)

- for i = l1 to u1 ' loop for strips

- for j = l2 to u2 ' loop for pixels

- debug.print(dBuf(i, j)

- next j : next i

- Sub AllocFrame(buffer as Variant)

- AllocFrame takes an empty variant and constructs a Visual Basic array of the correct size and data type to hold one frame of data from the document. Normally it is not necessary to use this function since GetFrame and GetStrip will perform an AllocFrame if the buffer Variant supplied to them is not set up yet.

- Example:

- Dim dx as WINX32Lib.DocFile

- Dim dBuf as Variant

- dx.Open("d:\data\circuit.spe")

- dx.AllocFrame(dBuf)

- '

- ' later in the program...

- '

- dx.PutFrame 1, dBuf

- Function GetParam(Param as enum DM_CMD, result as integer) as Any

- Function SetParam((Param as enum DM_CMD, Value as Any) as integer

- Allows setting and getting document parameters (strips, frames, etc.). SetParam returns non-zero on error. Note that not all parameters may be set after a document is created. For instance, once created the DM_XDIM and DM_YDIM parameters can not be set, since this would require changing the size of the file on disk.

- WinX/32 will try to convert the parameter to the correct type corresponding to the particular DM_CMD. If this cannot be done, Visual Basic reports a 'Type Mismatch' error. For more information see the discussion of Calling Syntax under Things to Note.

- Example:

- Dim dx as WINX32Lib.DocFile

- ydim = dx.GetParam(DM_XDIM) ‘ find the pixels per track

- dx.SetParam(DM_YLABEL, “Watts”) ‘ set the y label to “Watts”

- Function GetCalibration as CalibObj

- Function SetCalibration(Calib as CalibObj) as Boolean

- Allows setting and getting the X axis calibration information in the docfile. The information is passed as a CalibObj object, which is described below.

- Example:

- Dim dx as new Docfile

- Dim cx as CalibObj

- dx.Open(“Circuit.spe”)

- Set cx = dx.GetCalibration

- cx.DisplayUnits = XW_WAVELENGTH ‘ Change display units only

- bRes = dx.SetCalibration(cx)

- Sub Update()

- Redraws all frames for this doc.

- Function Save() as Boolean

- Saves the data in the document with the current filename.

- Function SaveAs(Name as String, nType as long) as Boolean

- Saves the data in the document with a new filename to a new file type. Use this function to convert a file to a new data type. The valid values for nType are: SPE  = 1,  TIFF = 2, TIF8 = 3.

- Function Close() as Boolean

Closes all windows associated with this doc and closes the file. If the doc is unsaved, the user will be prompted to save it before closing. The return value is 1 if the document was open before the save, or zero otherwise.

# DocFiles functions

The DocFiles object is a collection object used to refer to all of the open doc files. To use these functions, the user program must create an object of type DocFiles and call the functions through the object.

- Function GetAt(index as Integer) as DocFile

- Returns the open DocFile, if any, corresponding to the zero-based index. The index value must be less than the number of open documents, or else the returned object will be empty.

- Function Count() as Integer

- Returns the number of open DocFiles.

- Sub CloseAll()

Closes all of the open DocFiles.

# Doc Window functions

To use these functions, the user program must create an object of type DocWindow and call the functions through the object. A DocWindow object has these properties:

Properties:

| Name | Native type | Description |
| --- | --- | --- |
| Title | String | The title in the window (including the dimensions) |
| Visible | Boolean | The state of the shown/hidden window |

- Sub SetPosition(top as integer, left as integer, bottom as integer, right as integer)

- Sub GetPosition(top as integer, left as integer, bottom as integer, right as integer)

- Sets or gets the position of this window. Returns 0 on error.

- Sub Update()

- Redraws this frame.

- Function GetParam(Param as enum DI_PARAM, result as integer) as Any

- Function SetParam((Param as enum DI_PARAM, Value as Any) as integer

- Allows setting and getting DocWindow parameters (axes, cross sections, etc.). SetParam returns non-zero on error. It is recommended to call Update after calling SetParam. Not all parameters can be set after the window is created.

- For more information, see GetParam function in the DocFile section, and also the discussion of Calling Syntax under Things to Note.

- Function GetDocument() as DocFile

- Returns a pointer to an object of type DocFile, which you may use to call the DocFile interface functions for the document attached to this window.

- Function NewWindow([winType as PICTURE_TYPE]) as DocWindow

- Returns a pointer to a new Doc Window object representing a new window with the same document.

- Example:

- Dim dx as WINX32Lib.DocFile

- Dim tWin1 As WINX32Lib.DocWindow

- Dim tWin2 As WINX32Lib.DocWindow

- Dim r1, r2 As PICTURE_TYPE

- dx.Open("d:\data\circuit.spe")

- Set tWin1 = dx.GetWindow

- Set tWin2 = tWin1.NewWindow(WINX_GRAPH)

- Sub Next

- Sub Prev

- Sub First

- Sub Last

- Selects the strip or frame displayed by the window. For images, these subroutines control the displayed frame; for single curve graphs, they control the strip number; and for 3-D graphs, they control the strip number if the parameter DI_BPLOTFRAME is set to 0 or FALSE; otherwise, they control the frame number.

- Sub Update()

- Redraws this window.

- Sub Close()

- Closes this window.

- Sub ZoomIn

- Sub ZoomOut

- Sub UnZoom

- ZoomIn zooms in on the current window. If the ROI rectangle is being displayed in the window, the zoom is done to the ROI. If the ROI rectangle is not visible, then the zoom is done to increase the zoom factor by 10%.

- ZoomOut undoes the previous ZoomIn, if any. UnZoom undoes all the previous ZoomIns for this window.

- Function GetROI(index as Integer) as ROIRect

- Function SetROI(Rect as ROIRect) as Boolean

- Gets or sets the ROI rectangle for this window. Once the ROI rectangle is set, it can be displayed or hidden using the DI_BROIBOXON parameter in the SetParam function. The rectangle information is passed as an ROIRect object, which is described in that section.

- Function GetCursor as CursorObj

Gets the CursorObj for this window. You can use the CursorObj to read or set the position of the cross cursor in this window. CursorObj functions and properties are listed in the CursorObj section.

- Function GetCalibration as CalibObj

- Function SetCalibration(Calib as CalibObj) as Boolean

- Allows setting and getting the X axis calibration information for the window. The information is passed as a CalibObj object, which is described below. These functions are the same as the GetCalibration and SetCalibration functions for DocFile, except that they operate on the calibration of a window instead (the window calibration can be different from the file calibration, and different windows of the same data file can have different calibrations.)

- Function Activate() As Boolean

- This function makes this window the active window. The function returns True if successful, otherwise it returns False.

- Sub MoveCursor(X As Double, Y As Double)

This function sets to the cursor in the window to the coordinates (x,y).

- Sub SetROIPos(top As Double, left As Double, bottom As Double, right As Double)

This function draws an ROI rectangle in the window whose top-left point is (left, top) and whose bottom-right point is (right, bottom).

# Doc Windows functions

The DocWindows object is a collection object which refers to all of the open DocWindows. DocWindow properties referred to by SetParam or GetParam are the defaults for the WinX/32 program. To use these functions, the user program must create an object of type DocWindows and call the functions through the object.

- Function GetAt(index as Integer) as DocWindow

- Returns the open DocWindow, if any, corresponding to the zero-based. index. The index value must be less than the number of open DocWindows, or else the returned object will be empty.

- Function Count() as Integer

- Returns the number of open DocWindows.

- Sub CloseAll()

- Closes all of the open DocWindows.

- Sub Tile(BOOL bHorizontal)

- Sub Cascade();

- Arranges all of the open DocWindows.

- Function GetParam(Param as enum DI_PARAM, result as integer) as Any

Function SetParam(Param as enum DI_PARAM, Value as Any) as integer

Allows setting and getting DocWindow default parameters (axes, cross sections, etc.). SetParam returns non-zero on error. For more information see GetParam function in the DocFile section, and also the discussion of Calling Syntax under Things to Note.

- Function GetCalibration as CalibObj

- Function SetCalibration(Calib as CalibObj) as Boolean

- Allows setting and getting the default X axis calibration information for the program. The information is passed as a CalibObj object, which is described below. These functions are the same as the GetCalibration and SetCalibration functions for DocWindow, except that they operate on the default calibration of the program, rather than the calibration for a specific window

- Function GetActive() As DocWindow

- This function returns the active window if it’s a data window, otherwise it returns Nothing.

# Print Window functions

To use these functions, the user program must create an object of type PrintWindow and call the functions through that object. The PrintWindow appears when the user program first calls the Print or Clear methods of the object. A PrintWindow object has these properties:

Properties:

| Name | Native type | Description |
| --- | --- | --- |
| Title | String | The title in the window |

- Function OpenFile(Name as String) as Boolean

- Opens a print window file from disk. The file must already exist. The function returns False if not successful.

- Function Print(xpos as integer, line as integer, color as long, ...)

- Print to the print window at a character position xpos on a given line using a given color. Any number of items of any type can be printed in one call. Returns the number of characters printed. The color parameter takes a long hex number in the format &H00RRGGBB, where RR, GG, and BB can be in the range of 00 to FF.

Example:

Dim PWin as New WINX32Lib.PrintWindow

bRes = PWin.Print(0, 1, 0, "This is", text")

bRes = PWin.Print(5, 2, &HFF, "This is", " more", " text")

bRes = PWin.Print(10, 3, &HFF00, "This is", " line", 3)

- Function SaveAs(Name as String) as Boolean

- Saves a print window to a file on disk. If the file already exists, it is overwritten without warning. The function returns False if not successful. The file name should include an extension of “.PRN” in order for the file to be recognized by WinX/32 and Windows as a WinX/32 print window file.

- Function SetFont(FaceName as String, PointSize as Long) as Boolean

- This function sets the font for subsequently printed text in the print window. FaceName should be the name of an existing font on the system; if the font does not exist, Windows will try to find as close a match as possible. PointSize gives the size of the font in points. The function returns False if not successful

- Function SetPosition(top as integer, left as integer) as Boolean

- Set the position of the print window with respect to the screen (can exceed the boundaries of the app window). Returns False on error.

- Function SetSize(width as integer, height as integer) as Boolean

- Set the position of the print window with respect to the screen (can exceed the boundaries of the app window). Returns 0 on error.

- Sub SysCmd(Cmd as Long, Param as Long) as Long

This is a function for more advanced programmers. It allows a VB program to send a WM_SYSCOMMAND message to the print window. The WM_SYSCOMMAND message takes a parameter which causes the window to perform one of the actions in the “system menu”, the menu which appears when you click the box in the upper left corner of the window. The actions which a PrintWindow can perform in response to a WM_SYSCOMMAND message include, but are not limited to, those in the following table:

| Cmd | Value | Action performed |
| --- | --- | --- |
| SC_SIZE | &HF000 | Starts the keyboard-driven sizing of a window |
| SC_MOVE | &HF010 | Starts the keyboard-driven moving of a window |
| SC_MINIMIZE | &HF020 | Minimizes the window |
| SC_MAXIMIZE | &HF030 | Maximizes the window |
| SC_NEXTWINDOW | &HF040 | Selects the next open window |
| SC_PREVWINDOW | &HF050 | Selects the previous open window |
| SC_CLOSE | &HF060 | Closes the window |
| SC_VSCROLL | &HF070 | Causes the window to scoll vertically |
| SC_HSCROLL | &HF080 | Causes the window to scoll horizontally |
| SC_ARRANGE | &HF110 | Arranges the icons of minimized windows |
| SC_RESTORE | &HF120 | Restores the window from the minimized or maximized states |

The value for Param will depend on the value for Cmd; for all of the Cmds listed, it can be zero.

SysCmd returns a value which depends on the Cmd; for all of the Cmds listed, it will be zero if successful.

- Sub Clear()

- Clears the contents of the print window.

- Sub Close()

Closes the print window.

- Function GetRichEditWnd()as Long

This function is for advanced programmers only. It returns the handle of the window’s RichEditCtrl. The RichEditCtrl is a Windows object; PrintWindows use it to display text. By using the handle of the control, the programmer can send messages directly to the control in order to achieve text effects not supported directly by the PrintWindow.

- Sub Open()

This function opens the window.

# Print Windows functions

The PrintWindows object is a collection object, which refers to all of the open PrintWindows. To use these functions, the user program must create an object of type PrintWindows and call the functions through the object.

- Function GetAt(index as Integer) as PrintWindow

- Returns the open PrintWindow, if any, corresponding to the zero-based index. The index value must be less than the number of open PrintWindows, or else the returned object will be empty.

- Function Count() as Integer

- Returns the number of open PrintWindows.

- Sub CloseAll()

Closes all of the open PrintWindows.

# Experiment Setup functions

To use these functions, the user program must create an object of type ExpSetup and call the functions through that object.

- Function GetParam(Param as enum EXP_CMD, result as integer) as Any

- Function SetParam(Param as enum EXP_CMD, Value as Any) as integer

- Get and set experiment parameters (see the EXP_CMD parameters list for details).

- For more information see GetParam function in the DocFile section, and also the discussion of Calling Syntax under Things to Note.

- Function GetDocument() as Object

- Returns a pointer to an object of type DocFile. If the returned value is not empty, it is the data file associated with the current or last data collection. An empty return value indicates that either no data collection has run since the program started, or that the data file from the data collection  was closed. You can use the returned object to call the DocFile interface functions for the document.

- Example:

- Dim ex As New WINX32Lib.ExpSetup

- Dim dx As WINX32Lib.DocFile ‘ leave doc empty, WinX will supply one

- bRes = ex.Start(dx) ‘ will acquire one set and stop

- set dx = ex.GetDocument

- Function Start(File as DocFile) as Boolean

- Function StartFocus(File as DocFile) as Boolean

- Begin collecting data. Start causes all of the frames in the experiment to be acquired; then the acquisition stops. StartFocus acquires the first frame in the experiment repeatedly, until the Stop function is called (or until the WinX user stops the acquisition). If the DocFile object supplied is empty, WinX32 will create a new document and give it the default name for new documents, and the new data will be put into this document. If the DocFile object supplied is not empty, then the new data will be put into it. If an empty DocFile object is supplied, then on return it will be set to the new data file. The return value indicates whether data collection was successfully started.

- Example:

- Dim ex As New WINX32Lib.ExpSetup

- Dim dx As WINX32Lib.DocFile ‘ leave doc empty, WinX will supply one

- bRes = ex.Start(dx) ‘ will acquire one set and stop

- Function Stop() as Boolean

- Stop collecting data. The return value indicates whether data collection was in progress before the function was called.

- Function IsAvail(Param as enum EXP_CMD, optional Range as ValidRange) as Boolean

- Determines whether a feature corresponding to one of the parameters is available. You can call this function to find if the value returned from a given GetParam parameter will be valid, or to find whether the parameter value is able to be set via SetParam and what the range of valid values are.

- Example:

- Dim ex as New ExpSetup

- Dim vR as ValidRange

- bRes = ex.IsAvail(EXP_ADC_TYPE, vR) ‘ Can we set the ADC type?

- If bRes And vR.CurrentValue <> 2 And vR.MinValue <= 2 And vR.MaxValue >= 2 Then

- bRes = ex.SetParam(EXP_ADC_TYPE, 2) ‘ Yes

- End If

- Function GetROI(index as Integer) as ROIRect

Function SetROI(Rect as ROIRect) as Boolean

Gets or sets a detector ROI. ROIs determine how pixels are read from the detector. The value of index must be 1 or greater. The rectangle information is passed as an ROIRect object, which is described below. You must call SetParam with the EXP_USEROI parameter to control whether the ROIs are used. You should call ClearROIs before setting up one or more ROIs with SetROI.

The GetROI function can return an “empty” ROIRect object if the WinX software has no ROIs set up. If you try to use the “empty” ROIRect object, an error will occur. There are two ways to deal with this:

- Use an OnError Goto statement to catch errors in your program. In the OnError Goto handler, look at the global VB varible Err. If the Err variable indicates an “empty” object (if its value is 91), then you know that there were no ROIs set up.

- Declare the ROIRect variable in your program with “New”. A “New” ROIRect object has the Left, Right, Top and Bottom values set to 0, which are not valid for a “real” ROIRect. If, after calling GetROI, your ROIRect’s coordinates are all equal to zero, you can assume that there are no ROIs set up in WinX.

Example:

Dim ex as New ExpSetup, rx as New ROIRect

Set rx = ex.GetROI

if rx.Right <> 0 and rx.Left <> 0 then

points = rx.Right – rx.Left

End If

Example:

Dim ex as New ExpSetup, rx as ROIRect

OnError Goto NoROI

Set rx = ex.GetROI

if rx.Right <> 0 and rx.Left <> 0 then

points = rx.Right – rx.Left

End If

Finish:

OnError Goto 0

Goto NextThing

NoROI:

If Err = 91 then Points = 100 : Resume Finish

Else

OnError Goto 0

Resume

Endif

NextThing:

- Sub ClearROIs()

Erases all of the detector ROIs. This forces the detector to use the entire array for the next data collection. You can call SetParam with the EXP_USEROI parameter to control whether the ROIs are used without erasing them. You should call ClearROIs before setting up one or more ROIs with SetROI.

- Function Load as Boolean

Function Save as Boolean

These functions allow you to save and restore the entire experiment setup in a disk file. The name of the file is set by a previous call to the SetParam function with the EXP_SETUPNAME parameter. The return value is False if the operation did not succeed, otherwise it is True.

- Function WaitForExperiment as Boolean

This function waits for any pending data collection to complete. It is provided for programmers who do not wish to write DoEvents loops in their own code. The Start and StartFocus methods of ExpSetup return control almost immediately to the calling program. Often, a simple VB program that collects data will want to wait for the data collection to finish before proceeding. To do so, simply call this function:

Example:

Dim ex as New ExpSetup

Dim dx as DocFile

bRes = ex.Start(dx)

If (bRes) Then

BRes = ex.WaitForExperiment

End If

If you want your program to be able to do some processing while data collection is proceeding, you will have to write a loop using the VB DoEvents keyword (see the example on page 9.)

If the return value is False, then either there was no controller or the program was stopped before data collection completed.

These objects are used to pass data between the user’s program and WinX/32:

# ValidRange object

The ValidRange object is returned from the IsAvail function of ExpSetup. It has these properties:

MinValue 	– the minimum legal value for this parameter

MaxValue 	– the maximum legal value for this parameter

DefaultValue 	– the default value for this parameter (not always valid)

CurrentValue 	– the current value of this parameter

Increment 	– the minimum increment value for this parameter

ReadOrWrite 	– 1 = GetParam only, 2 = Get or Set

# CalibObj object

The CalibObj object is used by the GetCalib and SetCalib functions of DocFile. It has these properties:

Properties:

| Name | Native type | Description |
| --- | --- | --- |
| DisplayUnit | XCALIBUNIT | Selected calibration unit for Display |
| Input Unit | XCALIBUNIT | Selected calibration unit for the CalibVal array |
| Order | BYTE (1-5) | Order of polynomial for calibration |
| PairCount | Integer | Number of valid pairs of PixelLoc and CalibVal |
| PixelLoc(10) | Double | Array of locations specified in pixels |
| CalibVal(10) | Double | Array of calibration values at PixelLoc positions |
| PolyCoeffs(5) | Double | Array of polynomial coefficients for calibration |
| LaserPosition | Double | Laser wavenumber (valid for relative wavenumber units only) |
| Label | String | Calibration units label |
| UsageFlag | BYTE | Flag to use calibration; Auto-Spectro = 1, Manual = 2 or OFF = 3 |

Where XCALIBUNIT is one of:

XW_SYSTEM	Relative system units like channels, strips, frames

XW_DATA	Absolute units including offset and linear factor

XW_WAVELENGTH	Polynomial calibration ofwavelength

XW_ABSWAVENUM	Wavenumber units

XW_RELWAVENUM	Relative wavenumber units

XW_EVOLTS	Electron volts

The combination of DisplayUnits and InputUnits allows the programmer to supply calibration values (in the CalibVal array) in one unit, such as wavelength, while at the same time requiring the program to display the data (by default) in another unit, such as wavenumbers. The display units for a particular window may also be changed with the SetParam function of DocWindow.

The CalibObj object also has the following methods:

Function Lambda(Pixel as integer) as double - returns the wavelength for a given pixel.

Function Pixel(Lambda as double) as integer - returns the pixel for a given wavelength.

# DocInfo object

The DocInfo object supplies information to the OpenNew method of DocFile. It has the following properties:

| X | integer | Number of pixels in each strip of the new document |
| --- | --- | --- |
| Y | integer | Number of strips in each frame of the new document |
| Z | integer | Number of new frames in the new document |
| DataType | DATATYPE | Data type for the new document |
| Name | String | The name for the new document |
| bShowWindow | Boolean | True if the document should be shown in a window |
| bAppend | Boolean | True if the document should be opened in append mode |
| FileType | docType | Type of file (SPE, TIF, etc.) |

Here is an example of using the DocInfo object with OpenNew to open a new data file:

Dim dx As New DocFile

Dim ix As New DocInfo ‘ supplies info about new document

ix.x = 200 ‘ 200 pixels per strip

ix.y = 2 ‘ 2 strips

ix.Z = 1 ‘ 1 frame

ix.dataType = X_FLOAT ‘ floating point data

ix.bShowWindow = True ‘ display the file in a window

ix.bAppend = False ‘ overwrite; don’t append to existing file

ix.Name = "newdata" ‘ use this for the name (don’t supply extension)

ix.FileType = dt_SPE

bRes = dx.OpenNew("", ix) ‘ create the file

DocInfo defaults to bAppend = False, bShowWindow = True. If no name is supplied, the “Untitled” string is used. For example:

Dim dx As New DocFile

Dim ix As New DocInfo ‘ supplies info about new document

ix.x = 200 ‘ 200 pixels per strip

ix.y = 2 ‘ 2 strips

ix.Z = 1 ‘ 1 frame

ix.dataType = X_FLOAT ‘ floating point data

bRes = dx.OpenNew("", ix) ‘ create the file “untitled” with overwrite, show window

ROIRect object

The ROIRect object is used by the GetROI and SetROI functions of DocWindow and ExpSetup. It has the following properties:

Properties

Top 	– the X coordinate of the top left corner of the rectangle

Left 	– the Y coordinate of the top left corner of the rectangle

Bottom 	– the Y coordinate of the bottom right corner of the rectangle

Right 	– the X coordinate of the bottom right corner of the rectangle

XUnits 	– the units of the X coordinates

YUnits 	– the units of the Y coordinates

XGroup 	– the “group size” used in the X direction (only significant for ExpSetup ROIs)

YGroup 	– the “group size” used in the Y direction (only significant for ExpSetup ROIs)

- Sub Get(top As Double, left As Double, bottom As Double, right As Double,   xg As Long, yg As Long)

- Sub Set(top As Double, left As Double, bottom As Double, right As Double, xg As Long, yg As Long)

- Use these functions to get or set the ROI at once, without having to set each property.

# CursorObj object

The CursorObj object controls the cursor in a DocWindow. Call the GetCursor function of DocWindow to get the CursorObj for the window. The CursorObj properties are always current; that is, you don't have to do multiple GetCursor calls in order to get the most recent values, and when you set a  property, the cursor in the window moves immediately. (This is not like the ROIRect object, for instance.) All of the properties are of data type Double:

Properties

XPos 	– the X coordinate (pixel or wavelength) of the cursor

YPos 	– the Y coordinate (strip) of the cursor

ZPos 	– the Z coordinate (frame) of the cursor

Intensity 	– the intensity value of the pixel under the cursor

Parameter lists

# DM_CMD (DocFile Parameters)

Note: These functions only affect values in the data file, not in the controller.

| Parameter Name | Description Native Type |
| --- | --- |
| DM_ACCUMS | Accumulations |
| DM_ADCBITADJUST | ADC bit adjust |
| DM_ADCRATE | ADC rate |
| DM_ADCRESOLUTION | ADC resolution |
| DM_ADCTYPE | ADC type (FAST or SLOW) |
| DM_ASYNEN | Set asynchronous or synchronous (1 async, 0 sync) |
| DM_BACKGROUNDAPPLIED | Background subtraction flag |
| DM_BACKSUBNAME | Background subtraction file name |
| DM_BLEMISHAPPLIED | Blemish flag |
| DM_BLEMISHFILE | Blemish file name |
| DM_CLEANS | Cleans at start of acquisition |
| DM_CLKSPD | Clock speed, (kinetics/frame transfer) |
| DM_CLKSPD_us | Vertical clock speed in micro-seconds |
| DM_CONTROLLERNUM | Controller number if multiple controller system |
| DM_CONTROLLERTYPE | 3 = ST130, 5 = ST138, 6 = PentaMax, 7 = MicroMax, 9=VICCD |
| DM_COSMICAPPLIED | Cosmic ray removal flag |
| DM_COSMICTHRESHOLD | Cosmic ray removal threshold |
| DM_COSMICTYPE | Type of cosmic array if applied |
| DM_DATATYPE | Data type |
| DM_DETTYPE | Type of detector (i.e. Kodak 1400) |
| DM_DLABEL | Data label string |
| DM_DOCTYPE | Internal document type (SPE, TIFF, etc.) |
| DM_EXPOSEC | Exposure in seconds |
| DM_FILEDATE | Date file was created |
| DM_FILENAME | Name of file on disk |
| DM_FILETITLE | Name of file in window |
| DM_FILEVERSION | Software Version |
| DM_FLATFIELDAPPLIED | Flat field flag |
| DM_FLATFIELDNAME | Flat field file name |
| DM_FRAMECOUNT | Frames in this file |
| DM_FRAMEINPROS | Number of current frame in process/storage |
| DM_FRAMESIZE | Size of one frame in bytes |
| DM_GAIN | Gain |
| DM_GEOMETRIC | Geometric transformation: rotate = 1, reverse = 2, flip = 4 |
| DM_GPIBADD | GPIB address |
| DM_GPIBCONTROLADD | GPIB controller address |
| DM_HOURS | Hour of experiment |
| DM_HWACCUMS | Hardware Accumulation flag |
| DM_INTERFACETYPE | Interface card |
| DM_INTLEVEL | Interrupt level |
| DM_INUSE | Increment or decrement in use count |
| DM_IOADD1 | IO base address |
| DM_KINTRIGMODE | Kinetics Trigger Mode |
| DM_LASTFRAMERDY | Number of last frame finished acquiring |
| DM_MIN | Minutes of experiment |
| DM_MINBLK | Minimum block grouping for skip before valid pixel |
| DM_NUMEXPREPEATS | Number of Times Experiment Repeated |
| DM_NUMEXPACCUMS | Number of Times Experiment Accumulated |
| DM_NUMFRAMES | Number of frames in file |
| DM_NUMMINBLK | Number of min blocks to do before doing geometric |
| DM_NUMROI | Number of ROIs in header. If 0, 1 is assumed |
| DM_NUMSKPPERCLN | Number of skips per clean |
| DM_PIXELSIZE | Size of each data point in bytes |
| DM_READOUTMODE | Readout FULL/KINETICS/FRAME XFER |
| DM_SCANCOUNT | Scans in this file |
| DM_SECS | Seconds of experiment |
| DM_SHUTTERCOMP | Shutter compensation |
| DM_SHUTTERTYPE | Shutter type |
| DM_STOREFLAG | Flag to store incoming data |
| DM_STORESYNC | Store synchronous flag |
| DM_SWMADE | Software package ID: WinView = 1, WinSpec = 2, MicroView = 3 |
| DM_SWVERSION | File header version |
| DM_TEMPFLAG | Flag = TRUE if temporary file |
| DM_TIMINGMODE | Timing mode of experiment |
| DM_USERCOMMENT1 | User comment field 1 |
| DM_USERCOMMENT2 | User comment field 2 |
| DM_USERCOMMENT3 | User comment field 3 |
| DM_USERCOMMENT4 | User comment field 4 |
| DM_USERCOMMENT5 | User comment field 5 |
| DM_USERINFO | User information area |
| DM_WINDOWSIZE | Window size (kinetics) |
| DM_xCAL_CALIB_COUNT | Number of pixel-value pairs in array |
| DM_xCAL_COEFFS | Polynomial coefficient array |
| DM_xCAL_CURR_UNIT | The currently selected scaling unit |
| DM_xCAL_FACTOR | Factor for ditto |
| DM_xCAL_FILE_LABEL | Label for calibration file |
| DM_xCAL_INPUT_UNIT | The units of the PXL_VAL |
| DM_xCAL_LASERLINE | Laser line (nm) for rel cm-1 |
| DM_xCAL_OFFSET | Offset for absolute data scaling |
| DM_xCAL_POLYNOM_ORDER | Order of calibration polynomial |
| DM_xCAL_POLYNOM_UNIT | The units of the COEFFS |
| DM_xCAL_PXL_POS | Pixel array |
| DM_xCAL_PXL_VAL | Wavelength value array |
| DM_xCAL_STRING | The label for the axis |
| DM_xCAL_VALID | Flag is cal is valid |
| DM_XDIM | X dimension of actual data stored in this file |
| DM_XDIMDET | X Dimension of detector physical |
| DM_XLABEL | X label string |
| DM_YDIM | Y dimension of actual data stored in this file |
| DM_YDIMDET | Y Dimension of detector physical |
| DM_YLABEL | Y label string |

# DI_PARAM (DocWindow Parameters)

| Parameter Name | Description |
| --- | --- |
| DI_BAUTOASPECT | Maintain fixed image aspect ratio |
| DI_BAUTOZOOM | Maintain fixed image zoom factor |
| DI_BCOLBARON | Show color bar pane |
| DI_BFULLRANGEX | Set X axis to full range |
| DI_BHIDLINE | Draw 3-D with hidden line or surface |
| DI_BHIDSURF | Draw 3-D with hidden surface |
| DI_BHXSECTON | X cross section on or off |
| DI_BINFOBARON | Info Bar on or off |
| DI_BINITAUTOAUTO | TRUE autoscale every live scan |
| DI_BINITAUTORANGE | TRUE ignore range settings in page, FALSE, use them |
| DI_BINITAUTOSCALE | TRUE do autoscale, FALSE use intensity settings in page. |
| DI_BMARK3D | Draw cursor curve in differend color |
| DI_BMARKCURV | Draw 3-D marker curves in different color |
| DI_BNEWDATA | Flag that data may have changed |
| DI_BPLOTFRAME | Flag to plot 3-D in MultiStripSameFrame mode |
| DI_BPRINTCOLOR | Flag to print in color |
| DI_BROIBOXON | Flag to show ROI box |
| DI_BTOOLBARON | Tool Bar on or off |
| DI_BUPDATE | Flag to update window contents |
| DI_BVXSECTON | Y cross section on or off |
| DI_BXAXLBLON | X axis label on or off |
| DI_BXAXON | X axis on or off |
| DI_BYAXLBLON | Y axis label on or off |
| DI_BYAXON | Y axis on or off |
| DI_BZENDACTIV | Adjust Z axis position interactively |
| DI_DATAINTENHI | Max data intensity for display |
| DI_DATAINTENLO | Max data intensity for display |
| DI_DATASTEP | Pixel, strip, or frame skip value |
| DI_DATAVAIL | Bounds of available data (in data units) |
| DI_DATAVIS | Bounds of visible data (in data units) |
| DI_DBRITE | Window brightness setting 0 – 100 |
| DI_DCONTR | Window contrast setting 0 – 100 |
| DI_LABELDAT | Label string for data |
| DI_LABELXAX | Label string for x-axis |
| DI_LABELYAX | Label string for y-axis |
| DI_NEEDAUTO | 0 = no autoscale, 1 = do normal autoscale, 2 = do 5%/95% |
| DI_NMARKSTEP | Number of curves between 3-D marker curves |
| DI_PALETTE_SHAPE | Linear, log, histogram, etc. |
| DI_PALETTE_TYPE | Gray scale or false color |
| DI_PICTURE_TYPE | Graph, 3-D or Image (Read-only) |
| DI_ZAXISOFFSETX | X component of Z axis offset |
| DI_ZAXISOFFSETY | Y component of Z axis offset |
| DI_ZOOMXY | Ratio of picture size to data size, x & y. |

# EXP_CMD (ExpSetup Parameters)	U = Upgrade, S = Special

| Parameter Name | Set | Get | Avail | Description | Hardware Support |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | ST130 | ST138 | ST133 | ST133 5 MHz | Max | SpectraMax | View | PentaMax | Mini-Cam | MSP-100 | VICCD |
| EXP_5X_GAIN_ENABLE |  |  |  | 5X Gain Enable for 1MHz-only ADCs. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ABSORBFILENAME |  |  |  | Name of the experiment absorbance source file. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ABSORBSTRIP |  |  |  | Absorbance strip number. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ACCESSNAME |  |  |  | Name of the experiment access pattern save file. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ACCUMS |  |  |  | Number of accumulations to perform. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ACQBACK_SHUTTER |  |  |  | Acquire Background Shutter Mode. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ACQFLAT_SHUTTER |  |  |  | Acquire Flat Field Shutter Mode. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ACTUAL_TEMP |  |  |  | Detector actual temperature. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ADC_OFFSET |  |  |  | ADC Offset (0-255) for head. Not 1:1 equivalent to counts. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ADC_TYPE |  |  |  | ADC speed: Fast (8) or Slow (9) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_AMP_DIALOG |  |  |  | Amplifier type: Low noise(0) or High capacity(1). Requires Smart detector head. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ANALOG_GAIN |  |  |  | Sets the gain to low (1) medium (2), or high (3). Requires Smart detector head. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_AUTOSAVE |  |  |  | Automatically save named files |  |  |  |  |  |  |  |  |  |  |  |
| EXP_BLEMISHFILENAME |  |  |  | Name of the experiment blemish source file |  |  |  |  |  |  |  |  |  |  |  |
| EXP_CACCUMS |  |  |  | Number of accumulations performed so far |  |  |  |  |  |  |  |  |  |  |  |
| EXP_CONT_CLNS |  |  |  | Enable (1) or disable (0) continuous cleans |  |  |  |  |  |  |  |  |  |  |  |
| EXP_CONTROLLER_ALIVE |  |  |  | True if controller on and connected to computer |  |  |  |  |  |  |  |  |  |  |  |
| EXP_CONTROLLER_VERSION |  |  |  | Valid ranges (?) of controller version |  |  |  |  |  |  |  |  |  |  |  |
| EXP_COSMICFILENAME |  |  |  | Name of the experiment cosmic source file |  |  |  |  |  |  |  |  |  |  |  |
| EXP_COSMICSENS |  |  |  | Cosmic sensitivity |  |  |  |  |  |  |  |  |  |  |  |
| EXP_CSEQUENTS |  |  |  | Number of sequential frames collected so far |  |  |  |  |  |  |  |  |  |  |  |
| EXP_CUSTOM_SHUTTER |  |  |  | Custom shutter compensation (0-30). |  |  |  |  |  |  |  |  |  |  |  |
| EXP_DARKNAME |  |  |  | Name of the experiment background source file |  |  |  |  |  |  |  |  |  |  |  |
| EXP_DATATYPE |  |  |  | Experiment data type (include type AUTO) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_DATFILENAME |  |  |  | Default base name for the data file |  |  |  |  |  |  |  |  |  |  |  |
| EXP_DELAY_TIME |  |  |  | Delay between acquisitions. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_DISABLE_LASER |  |  |  | Sets optional ST133 register 0x06 to disable laser |  |  |  S |  |  |  |  |  |  |  |  |
| EXP_DOABSORB |  |  |  | Absorbance method (include disabled) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_DOBLEMISH |  |  |  | Blemish flag |  |  |  |  |  |  |  |  |  |  |  |
| EXP_DOCOSMIC |  |  |  | Cosmic flag (temporal, spatial, or disabled) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_EDGE_TRIGGER |  |  |  | Trigger polarity: Positive (1) or negative (0). |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ETACTUAL |  |  |  | “Set” is same as EXP_EXPOSURE. “Get” gets actual programmed exposure time from controller. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_EVENT_HANDLE |  |  |  | Win32 event handle; event is signaled on every new frame. Requires PCI or ISA interface; timer mode not allowed. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_EXPOSURE |  |  |  | Exposure Time desired in seconds |  |  |  |  |  |  |  |  |  |  |  |
| EXP_FILEACCESS |  |  |  | Overwrite/Append on new data file |  |  |  |  |  |  |  |  |  |  |  |
| EXP_FILEINCCOUNT |  |  |  | File increment count |  |  |  |  |  |  |  |  |  |  |  |
| EXP_FILEINCENABLE |  |  |  | File increment enable |  |  |  |  |  |  |  |  |  |  |  |
| EXP_FLATFLDNAME |  |  |  | Name of the experiment flat field source file |  |  |  |  |  |  |  |  |  |  |  |
| EXP_FORCE_PP_INIT |  |  |  | Force an initialization of the controller hardware. Requires PCI or ISA interface. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_HD_CHECK |  |  |  | Returns if the detector head is smart (has NVRAM) and if so what type it is. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_HTCLK |  |  |  | Horizontal Clock register (1-16) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_INTENSIFER_GAIN |  |  |  | Intensifier gain (0-255) for ICCD-Max only. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_KINETICS_WINDOWSIZE |  |  |  | Window size (active area) for hardware kinetics systems. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_LOGIC_OUTPUT |  |  |  | Output of BNC in the back of the controller: For ST133 V3, selects Not Scan (0) or Shutter (1). PentaMax V5 can also be Not Ready (2), Logic 0 (3), Cleaning (4), Not Frame Transfer Image Shift (5), and Logic 1 (7). |  |  | V3 |  |  |  |  |  V5 |  |  |  |
| EXP_MAXTHRESHOLDVALUE |  |  |  | Maximum Threshold Value |  |  |  |  |  |  |  |  |  |  |  |
| EXP_MIN_BLOCK |  |  |  | Minimum block size to bin ‘skips’ close to the valid data |  |  |  |  |  |  |  |  |  |  |  |
| EXP_MINTHRESHOLDVALUE |  |  |  | Minimum Threshold Value |  |  |  |  |  |  |  |  |  |  |  |
| EXP_NEWWINDOW |  |  |  | Use new window for every run |  |  |  |  |  |  |  |  |  |  |  |
| EXP_NOP_ADC |  |  |  | NOP count to slow up transfer speed. Some computers need this to use an ISA interface card with 1MHz controllers. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_NUM_MIN_BLOCK |  |  |  | Minimum number of skip blocks to use. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_NUM_OF_STRIPS_PER_CLN |  |  |  | Availability and value of number of strips per clean |  |  |  |  |  |  |  |  |  |  |  |
| EXP_NUMACCUMSEXP |  |  |  | Number of Times To Accumulate Experiment |  |  |  |  |  |  |  |  |  |  |  |
| EXP_NUMBER_OF_CLEANS |  |  |  | Availability and number of hardware cleans. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_NUMREPEATSEXP |  |  |  | Number of times to repeat entire experiment |  |  |  |  |  |  |  |  |  |  |  |
| EXP_OVERWRITECONFIRM |  |  |  | Overwrite confirmation |  |  |  |  |  |  |  |  |  |  |  |
| EXP_READOUT_TIME |  |  |  | Readout time of frame. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_RET_SRCCMP |  |  |  | Select whether to return source comp with data. Source comp uses one entire column. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_ROICOUNT |  |  |  | Gets number of ROIs stored (read-only) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_RS170 |  |  |  | RS170 availability flag. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_RUNNING |  |  |  | Get controller running status |  |  |  |  |  |  |  |  |  |  |  |
| EXP_RUNNING_APP |  |  |  | Get online process running status |  |  |  |  |  |  |  |  |  |  |  |
| EXP_RUNNING_EXPERIMENT |  |  |  | 1 = Entire Experiment Running, 0 = Not Running |  |  |  |  |  |  |  |  |  |  |  |
| EXP_SAVEROPT |  |  |  | Screen Saver/Blanker options |  |  |  |  |  |  |  |  |  |  |  |
| EXP_SEQUENTS |  |  |  | Number of sequential frames |  |  |  |  |  |  |  |  |  |  |  |
| EXP_SETUPNAME |  |  |  | Name of the experiment setup save file |  |  |  |  |  |  |  |  |  |  |  |
| EXP_SHT_GATE_MODE |  |  |  | Shutter/gating mode for ICCD-Max systems: Gated (1), Shutter (2), and Safe Mode (0). |  |  |  |  |  |  |  |  |  |  |  |
| EXP_SHT_PREOPEN |  |  |  | Mode for shutter: Pre-opened (1, usually in external sync mode), or opened and closed each shot (0). |  |  |  |  |  |  |  |  |  |  |  |
| EXP_SHUTTER_COMP_TIME_MS |  |  |  | This returns the shutter compensation time in ms. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_SKIPFRAMES |  |  |  | This skips frames in the beginning of the experiment. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_STORE_STROBE |  |  |  | Whether “store strobe” should be set. (1 = Yes) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_TEMP_STATUS |  |  |  | Temperature locked (1) or unlocked (0) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_TEMPERATURE |  |  |  | Controller Set Temperature |  |  |  |  |  |  |  |  |  |  |  |
| EXP_THRESHOLD_DEMAND_UPDATE |  |  |  | Threshold Demand Update |  |  |  |  |  |  |  |  |  |  |  |
| EXP_TIMING_MODE |  |  |  | Timing mode: Freerun (1) or external sync (3) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_TTL_LINES |  |  |  | Reads/Writes TTL lines (8 bits in and out) controller rear panel. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_USEMAXTHRESHOLDING |  |  |  | Maximum Thresholding: Disabled (0) or Enabled (1) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_USEMINTHRESHOLDING |  |  |  | Minimum Thresholding: Disabled (0) or Enabled(1) |  |  |  |  |  |  |  |  |  |  |  |
| EXP_USEROI |  |  |  | Use Full chip / ROI flag |  |  |  |  |  |  |  |  |  |  |  |
| EXP_VIDEO_TYPE |  |  |  | Video type: NTSC (1) or Pal (2). Also requires hardware jumpers to be installed. |  |  |  |  |  |  |  |  |  |  |  |
| EXP_VTCLK |  |  |  | Vertical Clock register (1-16). |  |  |  |  |  |  |  |  |  |  |  |
| EXP_XDIM |  |  |  | Get data X dimension; total size of all ROIs |  |  |  |  |  |  |  |  |  |  |  |
| EXP_XDIMDET |  |  |  | Get detector X dimension |  |  |  |  |  |  |  |  |  |  |  |
| EXP_YDIM |  |  |  | Get data Y dimension; total size of all ROIs |  |  |  |  |  |  |  |  |  |  |  |
| EXP_YDIMDET |  |  |  | Get detector Y dimension |  |  |  |  |  |  |  |  |  |  |  |
| EXP_YTNAME |  |  |  | Name of the experiment YT data file |  |  |  |  |  |  |  |  |  |  |  |

Creating a Snap-In DLL

To create a Snap-In DLL, use the PI Snap-In Generator Wizard to generate a generic SnapIn; then use the Resource Editor Add-In to add bitmaps for the toolbar buttons and strings for the menu item. The PI Snap-In Generator Wizard should be available from the Add-In menu. If it is not, try going to the Add-In Manager and make sure that the check box next to the Snap-In Generator Wizard is checked.

To use the wizard, begin by closing all open projects (use the Remove Project item of the File menu if necessary.) Then select the Snap-In Generator Wizard from the Add-In menu. Skip the first screen, and go to the screen that requests the project name. Click on the arrow to browse for a folder, and enter the name of your Snap-In. Then click on Open, and click on Next until the Finish button is available; then click on Finish. The Wizard will create various source files, then compile them to produce a Snap-In DLL. When it is done, you will get a message box with an OK button. After you click on OK, you will be able to edit the Snap-In code to produce your own actions. You will have to do two things to make your snap-in show up on the WinX32 toolbar:

1. Add toolbar bitmaps and menu string resources to the project

2. Register the DLL

To add the toolbar bitmap and menu string, you will need the resource editor Add-In. It is not shipped with Visual Basic, but is available from Microsoft: look for the file reseditl.exe on ftp.microsoft.com/softlib. Run this file to install the Add-In on your computer, then use the Add-In manager in Visual Basic to enable it. Then the Resource Editor will be an item on the VB Tools menu.

Use the Resource Editor to add two bitmaps to the project. You will have to use Windows Paint or a similar program to create the bitmaps first. One should be 16 X 16 pixels, the other 24 X 24 pixels. Give the smaller bitmap an ID of 101, and the larger an ID of 102. Next add two strings to the resource file. The first one will contain the menu item text for your snap-in, and should have an ID of 101. The second will have the status line text for the snap-in and should have an ID of 102. If you wish to have a tooltip for your snap-in, you will have to be able to insert a CHR$(13) + CHR$(10) sequence into the string somehow - the tooltip text is separated from the status line text by these two characters.

To register the snap-in, open the Immediate Window in Visual Basic (the menu item is under View). Then type the single command Register (case-sensitive) and hit Enter.

Now you should have a new menu item under Tools in WinX32, and a new toolbar with at least one button. Selecting the menu item or the toolbar button runs the code in the PiSnapIn-_OnCommand function of the clsSnapIn. You can put code in that function to bring up a form or do whatever you need.

How to Convert a Visual Basic “Exe” project to be a “Snap-In”

- In the Project Properties window, change the Project type to “ActiveX DLL”

- Add a module called DllMain. To this module, add a Public subroutine called Main

- In the Project Properties window, change the startup object to Sub Main

- In the Project References window, make sure “Princeton Instruments’s PIXBM32” and “Princeton Instruments’ Snap-In Library” are checked.

- Add the PI-supplied class clsSnapIn to the project

- To put up your main form (omit if your program doesn’t display a form):

- In some module (DllMain will do) add a Public variable of type Form:

- Public theForm as Form

- In the OnCommand subroutine in clsSnapIn, display the form as follows:

- Assign your form to the form variable

- Show the form in modal state

- Assign “nothing” to the form variable

Example:

Set theForm = frmMain

theForm.Show vbModal

Set theForm = Nothing

This approximates the behavior of an executable, where the form is loaded when the program starts, and is destroyed when the program ends. Your Snap-In DLL will be “running” as long as WinX/32 is running, but the toolbar icon or menu command will bring up the main form of the Snap-In; closing the form will destroy it.

- To run your DLL, you’ll need to add a menu string and (optionally) a toolbar bitmap. To add these things to a Visual Basic project, you’ll need the resource editor add-in. Install the add-in (you can get it from PI or from the Microsoft Web site) and enable it using the add-in manager. Then use it to add a resource file to your project and add the bitmap and string resources. Finally, in clsSnapIn’s GetMenuTextID and GetBitMapID functions, return the value of the ID of the string and bitmap, respectively.

- Compile your project. After it is compiled (successfully) open the Immediate Window and type the command “register” and hit Enter.

Comparison of Visual C++ and Visual Basic Interfaces to WinX/32

Advantages of Visual C++

- Visual C++ can access data in a data file directly through a pointer; Visual Basic has to copy it back and forth from an array.

- Visual C++ calls use what is called a vtable interface so they are faster than Visual Basic calls. This means that the function call itself is faster, not necessarily the action that happens during the function call.

- Visual C++ programs can do more complex operations more quickly than Visual Basic programs.

- Visual C++ DLL programs (“Snap-Ins”) can be debugged using the Visual C++ debugger; Visual Basic must be constructed as executables in order to use the Visual Basic debugger, but they can be converted to Snap-Ins if debugging is no longer necessary.

- Visual C++ DLL programs can “self-register”; that is, they can be installed on a machine which does not have Visual C++. Visual Basic DLL programs have to be registered from the command line of Visual Basic itself.

Advantages of Visual Basic

- Visual Basic prompts the user with information about every function call and parameter

- Visual Basic takes care of data types, initializing variables to zero, making sure pointers point to something, and a lot of other programming problems.

- Visual Basic makes it easier to design and code a user interface - just draw the form or dialog box with the controls you want, then click on the control to go to the code!

- Visual Basic programs compile faster than C programs

- Visual Basic is available to people using Excel, Word, etc.

- It is easier to use 3rd party controls ("ActiveX Controls") in Visual Basic than it is in C++

Index

1

16 X 16, 27

2

24 X 24, 27

3

3-D, 11, 21

3rd, 29

A

absolute, 20

Absorbance, 23, 24

Accumulation, 19

acquire, 14

acquired, 5, 14

acquires, 14

acquisition, 2, 5, 14, 19

acquisitions, 24

action, 2, 29

ActiveX, 28, 29

actual, 20, 23

actually, 5

ADC, 15, 19, 23, 25

address, 19

alive, 23

AllocFrame, 9

API, 6

appear, 2

appears, 12

application, 1, 6, 7

applications, 1

aquiring, 20

area, 20

arguments, 5

arise, 5

Arranges, 12

array, 4, 5, 8, 9, 16, 17, 19, 20, 29

aspect, 21

associated, 2, 9, 14

assume, 15

assumed, 20

async, 19

attached, 10

AUTO, 24

automatically, 1, 7, 8

Automation, 1, 6

autoscale, 21

Auto-Spectro, 17

available, 2, 5, 6, 15, 21, 25, 27, 29

available?, 25

axes, 10, 12

axis, 9, 11, 12, 20, 21

B

back, 3, 4, 29

background, 24

bar, 3, 6, 21

base, 19, 24

Basic, 1, 2, 3, 5, 6, 8, 9, 27, 28, 29

BB, 12

Beginning, 2

bit, 7, 19

bitmap, 27, 28

bitmaps, 27

Blemish, 19, 24

block, 20

blocks, 20, 25

bOn, 6

bottom, 2, 5, 6, 10, 18

boundaries, 13

Bounds, 21

box, 2, 3, 6, 21, 27, 29

Box, 6, 7

brightness, 21

briteness, 21

browser, 1, 2

bstr_t, 8, 17

buffer, 8, 9

button, 1, 2, 3, 5, 27

bytes, 19, 20

C

C, 1, 26, 29

C++, 1, 29

CalibObj, 9, 11, 12, 17

calibration, 9, 11, 12, 17, 20

CalibVal, 17

capabilities, 6

card, 19

care, 29

Cascade, 12

case, 5

causes, 14

change, 28

changed, 4, 17, 21

channels, 17

character, 12

characters, 5, 12, 27

check, 2, 4, 27

checked, 27, 28

choose, 1, 3

class, 28

clean, 20, 25

Cleans, 19

Clear, 12, 13

ClearROIs, 15, 16

Clears, 13

Click, 3, 4, 27

clicking, 1, 2, 5

clicks, 3

Clock, 19, 24, 26

Close, 9, 11, 13

CloseAll, 10, 12, 14

closed, 6, 14

CloseOpenDocs, 6

Closes, 6, 9, 10, 11, 12, 13, 14

closing, 9, 27, 28

cm-1, 20

code, 2, 3, 5, 16, 27, 29

coefficient, 20

coefficients, 17

COEFFS, 20

collecting, 5, 6, 14

collection, 2, 5, 10, 11, 14, 16

color, 12, 21

COM, 5, 7

combination, 17

command, 27, 28, 29

comment, 20

comp, 25

Comparison, 29

compensation, 20

compile, 1, 27, 29

completed, 16

complex, 29

computer, 27

confirmation, 25

consists, 1

constructed, 29

constructs, 9

contain, 27

contains, 6

contents, 3, 13, 21

contrast, 21

control, 1, 4, 11, 15, 16, 26, 29

controller, 7, 16, 19, 25

controlling, 6

controls, 2, 18, 29

convert, 9

converted, 1, 8, 29

coordinate, 18

coordinates, 6, 15, 18

copies, 8

copy, 8, 29

copying, 8

corner, 18

correct, 2, 8, 9

corresponding, 9, 10, 11, 14, 15

corresponds, 6

Cosmic, 19, 23, 24

count, 19, 24, 25

CountOpenDocs, 6

create, 2, 4, 6, 7, 10, 11, 12, 14, 18, 27

created, 2, 8, 9, 10, 19

creates, 5, 7, 8

cross, 10, 11, 12, 21

current, 7, 9, 11, 14, 17, 18, 19

currently, 20

CurrentValue, 15, 17

cursor, 11, 18, 21

curve, 11, 21

curves, 21

D

data, 2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 14, 16, 17, 18, 19, 20, 21, 24, 25, 26, 29

datatype, 7, 24

date, 8

deal, 15

debugged, 1, 29

debugger, 1, 29

debugging, 29

default, 7, 12, 14, 17

defaults, 2, 11, 18

DefaultValue, 17

defined, 4, 5

Delay, 24

described, 9, 11, 12, 15

Description, 17, 19, 21, 23

design, 29

details, 14

detector, 3, 7, 15, 16, 19, 20, 26

detectors, 6

determine, 7, 15

determines, 7

DI_BAUTOASPECT, 21

DI_BAUTOZOOM, 21

DI_BCOLBARON, 21

DI_BHIDLINE, 21

DI_BHIDSURF, 21

DI_BHXSECTON, 21

DI_BINITAUTORANGE, 21

DI_BINITAUTOSCALE, 21

DI_BMARK3D, 21

DI_BMARKCURV, 21

DI_BNEWDATA, 21

DI_BPLOTFRAME, 11, 21

DI_BPRINTCOLOR, 21

DI_BROIBOXON, 11, 21

DI_BUPDATE, 21

DI_BVXSECTON, 21

DI_BXAXLBLON, 21

DI_BXAXON, 21

DI_BYAXLBLON, 21

DI_BYAXON, 21

DI_BZENDACTIV, 21

DI_CMD, 21

DI_DATAINTENHI, 21

DI_DATAINTENLO, 21

DI_DATASTEP, 21

DI_DATAVAIL, 21

DI_DATAVIS, 21

DI_DBRITE, 21

DI_DCONTR, 21

DI_LABELDAT, 21

DI_LABELXAX, 21

DI_LABELYAX, 21

DI_NEEDAUTO, 21

DI_NMARKSTEP, 21

DI_PALETTE_SHAPE, 21

DI_PALETTE_TYPE, 21

DI_PARAM, 10, 11, 12

DI_PICTURE_TYPE, 21

DI_ZOOMXY, 21

dialog, 2, 29

did, 7, 16

differend, 21

different, 11, 21

dimension, 4, 5, 20, 26

dimensions, 4

direction, 18

directly, 29

disabled, 24, 26

disk, 7, 9, 16, 19

display, 2, 3, 4, 8, 9, 17, 18, 21, 28

displayed, 2, 7, 11

displaying, 7, 8

DisplayUnit, 17

DisplayUnits, 9, 17

ditto, 20

DLL, 1, 4, 27, 28, 29, 31

DLLs, 1

DM_ACCUMS, 19

DM_ADCBITADJUST, 19

DM_ADCRATE, 19

DM_ADCRESOLUTION, 19

DM_ADCTYPE, 19

DM_ASYNEN, 19

DM_BACKGROUNDAPPLIED, 19

DM_BACKSUBNAME, 19

DM_BLEMISHAPPLIED, 19

DM_BLEMISHFILE, 19

DM_CLEANS, 19

DM_CLKSPD, 19

DM_CLKSPD_us, 19

DM_CMD, 9, 19

DM_CONTROLLERNUM, 19

DM_CONTROLLERTYPE, 19

DM_COSMICAPPLIED, 19

DM_COSMICTHRESHOLD, 19

DM_COSMICTYPE, 19

DM_DATATYPE, 19

DM_DETTYPE, 19

DM_DLABEL, 19

DM_EXPOSEC, 19

DM_FILEDATE, 19

DM_FILENAME, 19

DM_FILETITLE, 19

DM_FILEVERSION, 19

DM_FLATFIELDAPPLIED, 19

DM_FLATFIELDNAME, 19

DM_FRAMECOUNT, 19

DM_FRAMEINPROS, 19

DM_FRAMESIZE, 19

DM_GAIN, 19

DM_GEOMETRIC, 19

DM_GPIBADD, 19

DM_GPIBCONTROLADD, 19

DM_HOURS, 19

DM_HWACCUMS, 19

DM_INTERFACETYPE, 19

DM_INTLEVEL, 19

DM_INUSE, 19

DM_IOADD1, 19

DM_KINTRIGMODE, 19

DM_LASTFRAMERDY, 8, 20

DM_MIN, 20

DM_MINBLK, 20

DM_NUMFRAMES, 20

DM_NUMMINBLK, 20

DM_NUMROI, 20

DM_NUMSKPPERCLN, 20

DM_PIXELSIZE, 20

DM_READOUTMODE, 20

DM_SCANCOUNT, 20

DM_SECS, 20

DM_SHUTTERCOMP, 20

DM_SHUTTERTYPE, 20

DM_STOREFLAG, 20

DM_STORESYNC, 20

DM_SWMADE, 20

DM_SWVERSION, 20

DM_TEMPFLAG, 20

DM_TIMINGMODE, 20

DM_USERCOMMENT1, 20

DM_USERCOMMENT2, 20

DM_USERCOMMENT3, 20

DM_USERCOMMENT4, 20

DM_USERCOMMENT5, 20

DM_USERINFO, 20

DM_WINDOWSIZE, 20

DM_xCAL_CALIB_COUNT, 20

DM_xCAL_COEFFS, 20

DM_xCAL_CURR_UNIT, 20

DM_xCAL_FACTOR, 20

DM_xCAL_FILE_LABEL, 20

DM_xCAL_INPUT_UNIT, 20

DM_xCAL_LASERLINE, 20

DM_xCAL_OFFSET, 20

DM_xCAL_POLYNOM_ORDER, 20

DM_xCAL_POLYNOM_UNIT, 20

DM_xCAL_PXL_POS, 20

DM_xCAL_PXL_VAL, 20

DM_xCAL_STRING, 20

DM_xCAL_VALID, 20

DM_XDIM, 4, 5, 9, 20

DM_XDIMDET, 20

DM_XLABEL, 20

DM_YDIM, 4, 9, 20

DM_YDIMDET, 20

DM_YLABEL, 9, 20

doc, 9, 10, 14

DocFile, 2, 3, 4, 5, 7, 8, 9, 10, 11, 12, 14, 16, 17, 18, 19

DocFiles, 10

document, 4, 5, 6, 7, 8, 9, 10, 14, 17, 18, 19

documents, 6, 10, 14

DocWindow, 8, 10, 11, 12, 17, 18, 21

DocWindows, 11, 12

DoEvents, 3, 4, 5, 16

double, 7, 17

Draw, 21

during, 29

dynamic, 1

E

Each, 1, 6

easier, 29

easiest, 1

EasyDLL, 6

editor, 27, 28

effort, 1

empty, 8, 9, 10, 11, 14, 15

enabled, 26

Enter, 27, 28

entry, 1

enum, 9, 10, 11, 12, 14, 15

Erases, 16

erasing, 16

error, 5

error-handling, 5

errors, 15

events, 4

every, 25, 29

example, 2, 3, 4, 5, 16, 18

Excel, 1, 29

except, 11, 12

executable, 1, 2, 28

executables, 1, 29

exist, 7

existing, 18

EXP_ABSORBFILENAME, 23

EXP_ABSORBSTRIP, 23

EXP_ACQBACK_SHUTTER, 23

EXP_ACQFLAT_SHUTTER, 23

EXP_ACTUAL_TEMP, 23

EXP_ADC_OFFSET, 23

EXP_ADC_TYPE, 15, 23

EXP_BLEMISHFILENAME, 23

EXP_CMD, 14, 15, 23

EXP_CONT_CLNS, 23

EXP_CONTROLLER_ALIVE, 23

EXP_CONTROLLER_VERSION, 23

EXP_COSMICFILENAME, 23

EXP_COSMICSENS, 23

EXP_DARKNAME, 24

EXP_DATATYPE, 24

EXP_DATFILENAME, 24

EXP_DELAY_TIME, 24

EXP_DISABLE_LASER, 24

EXP_DOABSORB, 24

EXP_DOBLEMISH, 24

EXP_DOCOSMIC, 24

EXP_ETACTUAL, 24

EXP_EXPOSURE, 24

EXP_FILEINCCOUNT, 24

EXP_FILEINCENABLE, 24

EXP_FLATFLDNAME, 24

EXP_FORCE_PP_INIT, 24

EXP_HTCLK, 24

EXP_MIN_BLOCK, 25

EXP_NOP_ADC, 25

EXP_NUM_OF_STRIPS_PER_CLN, 25

EXP_NUMBER_OF_CLEANS, 25

EXP_OVERWRITECONFIRM, 25

EXP_READOUT_TIME, 25

EXP_RET_SRCCMP, 25

EXP_RS170, 25

EXP_RUNNING, 3, 4, 5, 25

EXP_SAVEROPT, 25

EXP_SEQUENTS, 26

EXP_SETUPNAME, 16, 26

EXP_SW_TRIGGER, 26

EXP_TIMING_MODE, 26

EXP_TTL_LINES, 26

EXP_USEROI, 15, 16, 26

EXP_VIDEO_TYPE, 26

EXP_VTCLK, 26

Experiment, 4, 14, 20, 24, 25

Exposure, 19, 24

ExpSetup, 3, 4, 5, 14, 15, 16, 17, 18

F

facility, 1

factor, 11, 17, 21

fairly, 1

FALSE, 11, 21

family, 1

FAST, 19

faster, 29

feature, 15

few, 3

FF, 12

field, 19, 20

file, 1, 4, 8, 9, 11, 14, 16, 18, 19, 20, 23, 24, 26, 27, 28, 29

filename, 9

files, 10, 23, 27

Finally, 3, 28

find, 4, 8, 9, 15

Finish, 15, 27

finished, 4, 20

first, 2, 3, 7, 8, 12, 14, 27

fixed, 21

Flag, 17, 20, 21

Flat, 19, 23

flatfield, 24

flip, 19

following, 2, 3, 4, 17, 18

follows, 6, 7, 28

forces, 16

form, 2, 3, 5, 27, 28, 29

format, 12

forth, 29

frame, 3, 4, 8, 9, 10, 11, 14, 17, 18, 19, 20, 21

frames, 5, 8, 9, 14, 17, 20, 26

from, 1, 2, 3, 4, 7, 8, 9, 11, 14, 15, 17, 27, 28, 29

Full, 26

function, 2, 5, 6, 7, 8, 9, 10, 11, 12, 14, 15, 16, 17, 18, 27, 29

functions, 1, 2, 5, 6, 7, 8, 10, 11, 12, 14, 16, 17, 18, 19, 28

G

Gain, 19

generate, 5, 27

generation, 1

generic, 27

Geometric, 19

GetAppWnd, 6

GetCalib, 9, 11, 12, 17

GetCalibration, 9, 11, 12

GetDocument, 10, 14

GetFrame, 3, 4, 5, 8, 9

GetMenuTextID, 28

GetParam, 3, 4, 5, 9, 10, 11, 12, 14, 15, 17

GetPixel, 8

GetROI, 11, 15, 18

GetStrip, 8, 9

GetWindow, 7, 8, 10

GPIB, 19

Graph, 21

Gray, 21

group size, 18

grouping, 20

H

handle, 6

handler, 15

happens, 2, 29

Hardware, 19

head, 23

header, 20

height, 13

help, 2

hex, 12

hidden, 11, 21

hide, 7

histogram, 21

hold, 9

horizontal, 24

Hour, 19

I

I/O, 19

icon, 2, 28

ID, 20, 27, 28

identifies, 6

IDocWindowPtr, 8

ignore, 21

illustrates, 5

image, 3, 4, 21

IN_BLOCK, 25

include, 24

including, 8, 17

incoming, 20

increase, 11

Increment, 17, 19

index, 10, 11, 14, 15

information, 1, 3, 8, 9, 10, 11, 12, 14, 15, 17, 20, 29

init, 24

initializing, 29

Input, 17

InputUnits, 17

installed, 1, 2, 29

instance, 2, 5, 6, 9, 18

Instead, 8

Instruments, 1, 2, 28

int, 17

integer, 7, 8, 9, 10, 12, 13, 14, 17

integrated, 1

Intensity, 18

interactively, 21

interface, 2, 6, 8, 10, 14, 29

interfaces, 1, 2

Interrupt, 19

IO, 19

IsAvail, 15, 17, 25, 26

item, 1, 2, 3, 27

items, 2, 12

K

kinetics, 19, 20

L

Label, 17, 20, 21

language, 1

languages, 1

Laser, 17, 20

LaserPosition, 17

launches, 5

legal, 17

level, 19

libraries, 1, 2, 6

library, 1, 2

limitations, 1

line, 3, 5, 12, 20, 21, 27, 29

linear, 17

lines, 3, 26

link, 1

list, 1, 5, 6, 14

listed, 11

lists, 19

little, 3

live, 1

loaded, 28

location, 1

locations, 17

log, 21

Long, 6, 7

longer, 29

look, 1, 3, 15, 27

looking, 3

lot, 29

lx, 6

M

machine, 29

Macro, 1

made, 20

main, 6, 7, 28

Maintain, 21

make, 2, 3, 27, 28

makes, 29

making, 29

Manager, 27

Manual, 17

many, 4

marker, 21

match, 8

Max, 21

Maximize, 6

maximum, 17

MaxValue, 15, 17

may, 1, 2, 5, 8, 9, 10, 17, 21

means, 1, 5, 29

member, 2, 8

ment, 17

menu, 1, 2, 3, 27, 28

message, 3, 6, 27

messages, 5

method, 8, 17, 24

MicroMax, 19

Microsoft, 1, 27, 28

MicroView, 20

minimum, 17

Minutes, 20

MinValue, 15, 17

Mode, 19, 23

more, 3, 8, 9, 10, 12, 14, 15, 16, 29

most, 18

mouse, 3

multiple, 5, 18, 19

MultiStripSameFrame, 21

must, 1, 2, 5, 6, 7, 8, 10, 11, 12, 14, 15, 29

N

name, 2, 5, 7, 8, 14, 16, 17, 18, 19, 24, 27

named, 23

Native, 17, 19

necessarily, 29

necessary, 9, 27, 29

need, 5, 27, 28

new, 1, 2, 3, 5, 6, 7, 8, 9, 10, 14, 17, 18, 24, 25, 27

newName, 7

NewWindow, 10, 11

next, 8, 16, 27

nm, 20

NNING_APP, 25

non-zero, 6, 9, 10, 12

normal, 1, 21

Normally, 9

Note, 4, 5, 9, 10, 12, 14, 19

Now, 3, 27

NTSC, 26

null, 7

number, 5, 6, 8, 10, 11, 12, 14, 17, 19, 23, 25

O

object, 2, 4, 5, 6, 7, 8, 9, 10, 11, 12, 14, 15, 17, 18, 28

objects, 5, 17

offset, 3, 4, 17

OK, 27

old, 4, 8

OnCommand, 27, 28

OnError, 15

open, 2, 6, 7, 8, 9, 10, 11, 12, 14, 18, 27, 28

opened, 7, 8, 17

Opens, 6

operations, 29

optional, 7, 15

options, 25

Order, 17, 20

outside, 1

Overview, 1

Overwrite, 24, 25

P

package, 20

page, 2, 5, 16, 21

PairCount, 17

pairs, 17, 20

pane, 21

param, 4, 11

parameter, 2, 5, 7, 8, 9, 11, 12, 15, 16, 17, 29

parameters, 2, 5, 7, 8, 9, 10, 12, 14, 15

part, 2

particular, 9, 17

party, 29

pass, 17

passed, 9, 11, 12, 15

PentaMax, 19

people, 29

per, 9, 18, 20, 25

perform, 9

physical, 20

picture, 21

PICTURE_TYPE, 10, 21

pixel, 8, 17, 18, 20

PixelLoc, 17

pixels, 8, 9, 15, 17, 18, 27

pixel-value, 20

plot, 21

point, 3, 4, 7, 8, 18, 20, 29

pointer, 7, 8, 10, 14, 29

pointers, 29

points, 1, 3, 15

polynomial, 17, 20

port, 24

position, 10, 11, 12, 13, 21

positions, 17

prevent, 7

previous, 1, 2, 11, 16

Princeton, 1, 2, 28

Print, 5, 12, 14

printed, 12

PrintWindow, 5, 12, 14

PrintWindows, 14

problems, 29

proceed, 5

process, 5, 19, 25

processing, 2, 6, 16

produce, 1, 27

produced, 1

program, 1, 2, 3, 5, 6, 7, 9, 10, 11, 12, 14, 15, 16, 17, 27, 28

programmer, 17

programmers, 16

Programming, 1

programs, 1, 4, 29

progress, 14

project, 1, 2, 27, 28

prompted, 9

prompts, 29

proper, 8

properly, 2

properties, 11, 17, 18

property, 18

provide, 1, 6

provided, 1, 16

providers, 1

provides, 1, 2, 5

PutFrame, 4, 8, 9

PutPixel, 8

PutStrip, 8

PXL_VAL, 20

Q

quickly, 1, 29

R

range, 12, 15, 21

rate, 19

ratio, 21

ray, 19

read, 11, 15

Read-only, 21

ReadOrWrite, 17

Readout, 20, 25

real, 15

recommended, 10

rectangle, 11, 15, 18

Redraws, 9, 10, 11

References, 1, 2, 28

referred, 11

register, 27, 28

registered, 1, 29

registry, 1

relative, 17

resource, 1, 27, 28

reverse, 19

ROI, 11, 15, 21, 26

ROIRect, 11, 15, 17, 18

ROIs, 15, 16, 18, 20, 25

rotate, 19

RS170, 25

run, 1, 2, 5, 14, 25, 28

running, 1, 3, 4, 6, 7, 25, 28

S

same, 8, 10, 11, 12, 17

save, 9, 16, 23, 26

Saves, 9

scale, 21

scaling, 20

screen, 7, 13, 27

second, 27

seconds, 7, 19

section, 4, 6, 10, 11, 12, 14, 21

sections, 10, 12

see, 2, 3, 4, 8, 9, 10, 12, 14, 16

select, 2, 27

Selected, 17

self-register, 29

SendMessage, 6

sensitivity, 23

separate, 5

separated, 27

sequential, 26

set, 1, 2, 5, 8, 9, 10, 11, 14, 15, 16, 18, 23

SetCalib, 17

SetCalibration, 9, 11, 12

SetParam, 5, 9, 10, 11, 12, 14, 15, 16, 17

SetPosition, 5, 10, 12

sets, 8, 11, 15

setting, 5, 9, 10, 11, 12, 15, 16, 21

settings, 21

setup, 4, 7, 16, 26

short, 2, 6

should, 2, 3, 5, 8, 15, 16, 17, 27

show, 1, 2, 4, 7, 18, 21, 27

ShowDemoBox, 2, 3, 6

ShowWindow, 6

Shutter, 20, 23

significant, 18

similar, 8, 27

simple, 6, 16

simply, 16

since, 9, 14

single, 7, 8, 11, 27

size, 7, 8, 9, 18, 20, 21

skip, 20, 21, 25

skips, 20, 25

SLOW, 19

slower, 1

smaller, 27

SnapIn, 27

Snap-In, 1, 4, 27, 28, 31

Snap-Ins, 1, 29

software, 1, 15, 20

some, 1, 2, 5, 16, 28

something, 2, 3, 5, 29

source, 23, 24, 25, 27

spatial, 24

specific, 12

specified, 8, 17

specify, 8

specifying, 2

speed, 1, 19, 25

ST130, 19

ST133, 23, 26

stand-alone, 1, 4

standard, 1, 2

Start, 2, 3, 4, 5, 14, 16

started, 3, 14

StartFocus, 5, 14, 16

starting, 7

starts, 2, 28

start-up, 2

state, 28

statement, 15

status, 25, 27

stop, 5, 14

stops, 14

storage, 19

store, 20

stored, 20, 25

stores, 1

string, 2, 6, 7, 18, 19, 20, 21, 27, 28

strip, 8, 11, 17, 18, 21, 23

strips, 8, 9, 17, 18, 25

structure, 8

subtraction, 19

succeeded, 5

successfully, 14, 28

such, 1, 5, 7, 8, 17

supplied, 7, 9, 14, 18

supply, 14, 17, 18

support, 1

supported, 6

supports, 1, 6

sure, 2, 3, 27, 28, 29

surface, 21

SW_SHOWMAXIMIZED, 6

sync, 19

synchronous, 19, 20

system, 17, 19

T

takes, 3, 9, 12, 29

task, 5

technical, 1

temperature, 23

temporal, 24

temporary, 20

text, 6, 12, 27

thread, 5

three, 1

threshold, 19

through, 4, 6, 7, 10, 11, 12, 14, 29

Tile, 12

time, 7, 17, 25

Timing, 20

tool, 27

toolbar, 1, 27, 28

Tools, 27

tooltip, 27

top, 2, 5, 6, 10, 12, 18

transfer, 19, 25

transformation, 19

Trigger, 19

TRUE, 6, 17, 20, 21

TTL, 26

type, 1, 2, 3, 5, 6, 7, 8, 9, 10, 11, 12, 14, 15, 17, 18, 19, 20, 24, 26, 27, 28

types, 29

U

under, 2, 3, 9, 10, 12, 14, 18, 27

undoes, 11

unit, 17, 20

units, 9, 17, 18, 20, 21

unsaved, 9

UnZoom, 11

Update, 4, 9, 10, 11, 26

urces, 27, 28

UsageFlag, 17

V

valid, 8, 9, 15, 17, 20

ValidRange, 15, 17

variable, 5, 8, 15, 28

variables, 29

VARIANT, 8, 9, 11

VARIANT_BOOL, 8, 9, 11, 17

variant_t, 11

various, 27

version, 6, 20

VICCD, 19

Visual, 1, 2, 3, 4, 5, 6, 8, 9, 27, 28, 29

Visual Basic, 1, 2, 3, 4, 5, 6, 8, 9, 10, 27, 28, 29

Visual C++, 1, 29

volts, 17

vtable, 29

W

Wavelength, 20

wavenumber, 17

wavenumbers, 17

width, 13

Win32, 6

window, 1, 2, 3, 5, 6, 7, 8, 10, 11, 12, 13, 17, 18, 19, 21, 28

Windows, 1, 5, 6, 11, 14, 27

Winspec, 1, 6

Winview, 1, 6

WinX, 1, 2, 3, 4, 5, 6, 7, 9, 11, 14, 15, 17, 28, 29

WinX32, 1, 2, 3, 6, 14, 27

Winx32App, 2, 3, 6

WINX32Lib, 8, 9, 10, 12, 14

Wizard, 1, 27

writable, 23

write, 1, 2, 5, 16

written, 1

X

x-axis, 21

XCALIBUNIT, 17

XGroup, 18

xLen, 7

xpos, 12

XUnits, 18

XW_ABSWAVENUM, 17

XW_DATA, 17

XW_EVOLTS, 17

XW_RELWAVENUM, 17

XW_SYSTEM, 17

XW_WAVELENGTH, 9, 17

Y

y-axis, 21

YGroup, 18

yLen, 7

YUnits, 18

Z

zero, 8, 9, 15, 29

zero-based, 10, 11, 14

zoom, 11, 21

ZoomIn, 11

ZoomOut, 11